Cleaning
1) Flowing water wash the instrument, get rid off blood or mucus
left ,use high pressure water gun flush inner tube .
2) Soak 1~10 minutes in multi enzyme solution
3) Flowing water wash , and use water gun clean inner tube .
4) Use air gun blow the instrument to clean white gauze to see
whether clean enough ,no other colors show in the gauze .
5) Dry totally .

Sterilization
Sterilization operation process please follow the medical institution
disinfection and supply center surgical instruments disinfection and
sterilization procedures .

Methods as below :
1) Ethylene oxid fumigation;
2) Formalin steam sterilization;
3) Low-temperature plasma sterilization
4) 2% glutaraldehyde solution for 10h.
5) Cidex immerse.
6) Autoclave via standard sterile.( Note : In autoclave standard time
and temperature set as follow, 115degrees around 30minutes , 121
degrees around 12 minutes ,132 degrees around 2 minutes . For our
instrument , time no more than 20minutes ,temperature no more than125 degrees .)Above methods recommend EO better , and autoclave recommend 132degrees around 2 minutes .